HIMMS and RSNA

IHE

Integrating the Healthcare Enterprise

Test Plan for Image Creator

2003-12-19

Marco Eichelberg1,  Klaus Kleber2, Jörg Riesmeier1

1	Kuratorium OFFIS e.V., Escherweg 2, 26121 Oldenburg, Germany

2	VISUS Technology Transfer GmbH, Philippstraße 5, 44803 Bochum, Germany

Contents

1	Introduction	3

2	Test Procedure for Image Displays	4

# Introduction

This document is intended for vendors participating in the IHE demonstration which will supply one or more Image Creator Actors as defined in the IHE Technical Framework.  In principle, it is also applicable to Acquisition Modality Actors.

Image Creators are systems that create additional images and/or Grayscale Softcopy Presentation States and transmit the data to an Image Archive.  They also make requests for the storage commitment of the Image Manager for the images and/or Presentation States previously transmitted and generate Performed Procedure Steps.  The Image Creator must be grouped with an Image Display. They are required to implement the following IHE transactions:

- “Query Image”, transaction 14 of the IHE Radiology Technical Framework.

- “Retrieve Image”, transaction 16 of the IHE Radiology Technical Framework.

- “Creator Images Stored”, transaction 18 of the IHE Radiology Technical Framework.

- “Creator Presentation State Stored”, transaction 19 of the IHE Radiology Technical Framework.

- “Creator Storage Commitment”, transaction 22 of the IHE Radiology Technical Framework.

They can optionally support the following IHE transactions:

- “Query Presentation States”, transaction 15 of the IHE Radiology Technical Framework.

- “Retrieve Presentation States”, transaction 17 of the IHE Radiology Technical Framework.

- “Creator Procedure Step in Progress”, transaction 20 of the IHE Radiology Technical Framework.

- “Creator Procedure Step Completed”, transaction 21 of the IHE Radiology Technical Framework.

If the Image Creator can be operated in different modes which exhibit a different behavior on the DICOM network protocol level, the tests should be run for each mode of operation.

# Test Procedure for Image Creators

The IHE MESA Tools contain, among other components, a Java-based GUI application called “DICOMscope”.  This application is used in the test procedure for Image Creators.  Installation instructions and a simple user manual for the DICOMscope application are provided in different documents with the IHE MESA distribution.  The following discussion assumes that the application has been successfully installed and that the user has a basic understanding of the application’s concepts as laid out in the user manual.  However, not all functions offered by the DICOMscope application are required or used in this test procedure.

DICOMscope contains a Query/Retrieve Server that implements the requirements of the IHE “Query Image”, “Retrieve Image”, “Query Presentation State”, “Retrieve Presentation State” transactions.  It should be noted that interoperability between an Image Creator and the DICOMscope testing Query/Retrieve Server does improve the probability for interoperability between the SCU and SCP of the specified provided by a different vendor, but is in no way a guarantee for interoperability.

Images and Presentation States created with any Image Creator should be validated with DICOMscope.  DICOMscope is a Storage Service Class Provider.  Every Image Creator can send images and/or Presentation States to DICOMscope where they can be visualized.  It should be noted that DICOMscope is not able to display color images and that Presentation States are only defined for grayscale images.

The Image Creator should be grouped with an Image Display.  Therefore, the Test Plan for the Image Display should also be checked.