Guidelines for Cross-Vendor DICOM Testing

Title:			Application Entities Interconnectivity

Association and CT Storage Services

Level_3 Test-Design Specification Example

Document ID:		1.2.840.113702.1.3.13.1

Version Date:		1/20/96

Application Entities Interconnectivity

Association and CT Storage Services

Level_3 Test-Design Specification Example

i)	Purpose

The purpose of this Test Design Specification document is to detail refinements of the test approach presented in the Test Plan for interconnectivity testing of the General Electric Advantage Review Workstation. The Test Design Specification  will identify the features to be tested by the design and its associated tests, identify procedures for executing the defined tests, and identify specific test cases.

This document is the Level_3 Test Design example document for the combined Association Services and CT Storage Service Class. Association Services and Storage are operationally inseparable for the Advantage Review Workstation SCP and therefore the Test Plan, Test Design and Test Case documents have been written to reflect this operational reality. DICOM interconnectivity testing of the Advantage Review Workstation does not require multiple separable Test Design documents, nor is it necessary to have a separable Test Procedure document. This document is intended to represent a specific Test Design and Procedure which accommodates all of the characteristics of the specified Application Entity, which complies with the DICOM Standard, ACR/NEMA Standards Publication No. PS 3.

ii)	Outline

A test-design specification document shall have the following structure.

1.	Test-design-specification identifier

2.	Features to be tested

3.	Approach refinements

4.	Test identification

5.	Feature pass/fail criteria

6.	Special Requirements

7. 	Procedure Steps

1.	Test-Design-Specification Identifier

This example Test Design Specification has the ID: 1.2.840.113702.1.3.13.1

2.	Features to be Tested

The features which are testable and the design of tests for those features is completely dependent on the access that the testers have to enter and observe information in the systems under test. While the Advantage Review Workstation is capable of being both a Service Class Provider and a Service Class User of the CT Store SOP, the cross-vendor testing situation defined in this document allows for testing only the Service Class Provider SOP capability of this system. The operational granularity and the access to logging information that is accessible to testers differs considerably from the idealized model presented in the Testing Guidelines.

The following features of the Local DICOM Import Application Entity as implemented in Software Revision 1.03 of the Advantage Review Workstation will be tested:

Basic TCP/IP connectivity

Basic Association Establishment and image transfer

SCP role of the CT Image Storage SOP Class

Association Abort

Behavior at Boundary Cases

Repeatability

3.	Approach Refinements

The Advantage Review Workstation provides reasonably detailed information that is useful for testing, but that information is presented in a window printout that is continually updated and partially overwritten and is apparently not logged.  The difficulty, therefore, is to catch the messages going by and to correctly interpret the sequence. From the associated Conformance Statement, and from all observations, the Advantage Review Workstation does not establish concurrent associations. Each message displayed to the operator is date and time stamped, but for simplicity the date/time information will not be included in sample messages described herein. As images are received in DICOM format, the Advantage Review Workstation takes each message in sequence and converts the image data to an internal format and stored in the local database. That conversion does proceed concurrent with receiving additional images, and messages indicating each conversion are also displayed to the window. Those messages will also be omitted from our testing description since they are outside the scope of DICOM interconnectivity testing. It should be noted, however, that the information output of the conversion process is extremely relevant to DICOM Interoperability, as defined in the Introduction and Overview Document of the Testing Guidelines.

The testing relevant messages displayed by the Advantage Review Workstation are:

1. Connection (date time) Starting

2. Connection (date time) Running DICOM:  Receiving Images

3. Connection (date time) Running  DICOM Import Received :  n Images from Application Entity Name.

4. Connection (date time) Exiting  DICOM Import Received :  n Images from Application Entity Name.

5. Connection (date time) Abort DICOM Import Received :  n Images from Application Entity Name.

4.	Test Identification

By analysis of the conformance statement of the GE Advantage Review Workstation the following interconnectivity tests have been designed:

| Feature | Test Case ID | Test Case Description |
| --- | --- | --- |
| Basic TCP/IP connectivity | 1.2.840.113702.1.3.14.1 | Validation of TCP configuration and connectivity |
| Basic Association Establishment and image transfer (SCP role) | 1.2.840.113702.1.3.15.1 | Single study |
| Association Establishment and Image Transfer - Multiple Study | 1.2.840.113702.1.3.16.1 | Select/Queue multiple cases; 3 different patients, 4 studies ( 2 studies of one patient) |
| Repeatability of Transfers | 1.2.840.113702.1.3.17.1 | Repeatability Test; Single patient, resend the study |
| Association Abort | 1.2.840.113702.1.3.18.1 | Abort; Start sending a long series of images and Abort |
| Boundary Behavior - Multiple Simultaneous Associations | 1.2.840.113702.1.3.19.1 | Queue up multiple cases and try to force simultaneous associations |
| Boundary Behavior - Resource boundary | 1.2.840.113702.1.3.20.1 | Full disk |

5.	Feature Pass/Fail Criteria

| Feature | Test Case ID | Pass Criteria |
| --- | --- | --- |
| Basic TCP/IP connectivity | 1.2.840.113702.1.3.14.1 | node is alive |
| Basic Association Establishment and image transfer (SCP role) | 1.2.840.113702.1.3.15.1 | DICOM Import Received : n Images from Application Entity Name |
| Association Establishment and Image Transfer - Multiple Study | 1.2.840.113702.1.3.16.1 | For each case - DICOM Import Received : n Images from Application Entity Name |
| Repeatability of Transfers | 1.2.840.113702.1.3.17.1 | DICOM Import Received : n Images from Application Entity Name or warning message indicating previously stored UID. |
| Association Abort | 1.2.840.113702.1.3.18.1 | DICOM transfer aborted error message |
| Boundary Behavior - Multiple Simultaneous Associations | 1.2.840.113702.1.3.19.1 | For each case - DICOM Import Received : n Images from Application Entity Name |
| Boundary Behavior - Resource boundary | 1.2.840.113702.1.3.20.1 | disk full error condition. |

6.	Special Requirements

The workstation must be physically connected to a local area network. Both TCP/IP and DICOM configuration must be established such that it is possible to ÒpingÓ the GE machine from another network device.  Domain and name services must be provided on the network so that a host name to IP address translation is supported for all systems under test.

7. Procedure Steps

7.1 Log Create a test log data sheet and record all system status responses.

7.2 Setup Configure TCP/IP communication. Add appropriate Application Entity Titles, IP addresses and 	port numbers to the database. Configure the workstation in configuration 1, no simultaneous associations.

7.3 Measure

1.	Single study

2.	Select/Queue multiple cases; 3 different patients, 4 studies ( 2 studies of one patient)

3.	Repeatability Test;  Single patient, resend the study

4.	Abort; Start sending a long series of images and Abort

5.	Queue up multiple cases and try to force simultaneous associations

6.	Fill the workstationÕs disks and attempt an image transfer.

7.4 Wrap up Compile Test Summary.

7.5 Contingencies Compile and submit Test Incident Reports. If remedies or work-arounds are possible, implement and retest.